Notation

- I = background-corrected intensity (CPS = counts per second)

- CPSbg = mean background intensity

- CPSmat = mean matrix intensity

- dt = dwell time (s)

- N = number of sweeps in the window

- S = sensitivity = ICAL/C

- C = certified mass fraction (RM)

- σ = 1σ standard uncertainty

- w = fractional position in a drift segment

Subscripts / indices

- i = the analyte being quantified (the element of interest)

- IS = the internal standard

- n = running index over all analytes (the summation index in total oxides)

- A, B = the two bracketing CAL anchors of a drift segment

- CAL = a calibration-standard measurement

- ,i (as in σS,counting,i) = the i-th CAL spot within a group

- ,counting / ,RM = the counting and reference-material components

- rel = relative to the internal standard (Srel=Si/SIS)

- source = the calibrator a "from-CAL" internal standard is inherited from

Derived quantities (grouping + total oxides)

- Smean= group-mean sensitivity (mean of the CALs in a group)

- MSWD = mean square weighted deviation (reduced χ² of the CALs about their mean)

- pn= preliminary mass fraction of element n=maxIn,0/Sn

- cn = oxide conversion factor of element n (0 for non-oxide analytes)

- D=closure sum=Σnpn·cn(over oxide-forming majors)

- T = target total = totaloxidewt·104µg/g

- λ = normalization constant = T/D

- δ_(in) = Kronecker delta (1 if n = i, else 0)

Step 1 — Counting SEM on the raw intensities

σbg=√CPSbg/dt·Nbg

σmat=√CPSmat/dt·Nmat

Explanation: the 1σ standard error of the mean intensity in the background and matrix windows, from Poisson counting statistics.

Reasoning: the instrument reports counts per second, so total registered counts in a window=CPS·dt·N. Poisson gives σ=√counts, so the relative counting uncertainty is 1/√CPS·dt·N. We use the standard error of the mean (not the sample standard deviation) because each concentration uses the mean over N sweeps. The dwell time is essential — omitting it under-estimates the counting term by √dt.

Step 2 — Background-corrected intensity SEM

I=CPSmat-CPSbg

σI=√σmat2+σbg2

Explanation: the net signal is background-subtracted; its uncertainty combines the two window SEMs.

Reasoning: for a subtraction, absolute uncertainties add in quadrature. Background and matrix are measured in separate time windows, so their counting errors are independent.

Step 3 — Per-CAL sensitivity uncertainty (two components)

S=ICAL/C

σS/S2=σI/I2+σC/C2

σS,counting=S·σI/I

σS,RM=S·σC/C

Explanation: the sensitivity is intensity over certified mass fraction; a division, so relative uncertainties add in quadrature. The two contributions are kept separate.

Reasoning: they behave differently downstream. Counting is independent per CAL spot (random). RM is systematic — the same certified value is used for every CAL of that material, so it is fully correlated. Merging them early would give the wrong result when averaging and interpolating, so they are carried as two components.

Step 4 — Group averaging of CALs

Smean=meanSi

counting (independent → shrinks with N):

σmean,counting=√ΣσS,counting,i2/N

RM (correlated → does NOT shrink):

σmean,RM=Smean·σC/C

overdispersion (MSWD):

MSWD=1/N-1·ΣSi-Smean2/σS,counting,i2

σmean,counting=σmean,counting·√max1,MSWD

Explanation: averaging repeated calibrations sharpens the random part but not the systematic part. The MSWD then asks whether the CALs actually agree within their counting uncertainties.

Reasoning: independent errors reduce as the standard error of the mean (~√N); a correlated systematic offset does not reduce at all. MSWD (reduced χ²) captures disagreement between CALs that pure counting statistics miss — if the CALs scatter more than their counting σ predicts, the counting term is inflated by √MSWD. RM is untouched (it causes no scatter within a group).

Step 5 — Drift interpolation to the unknown

counting (independent anchors → quadrature):

σcountingw=√1-w2·σA,counting2+w2·σB,counting2

RM (correlated anchors → linear):

σRMw=1-w·σA,RM+w·σB,RM

combined:

σSw=√σcountingw2+σRMw2

Explanation: the sensitivity applied to an unknown is a linear interpolation between the two bracketing CAL anchors A and B; this is its confidence band along the drift.

Reasoning: the two anchors' counting errors come from separate measurements → independent → quadrature (band narrows between anchors). The RM error is the same systematic offset at both anchors (same reference material) → fully correlated → interpolates linearly. Quadrature-ing the RM part would wrongly shrink it between calibrations.

Step 6 — Concentration propagation

6a. Direct internal standard

Ci=Ii/IIS·CIS/Srel,iSrel,i=Si/SIS

σC,i/Ci2=σI,i/Ii2+σI,IS/IIS2+σS,i/Si2+σS,IS/SIS2

IS column: σ=0

Explanation: the mass fraction is a pure product/quotient of the analyte and IS intensities and sensitivities.

Reasoning: for products/quotients, relative uncertainties add in quadrature, and each factor contributes whether it sits in numerator or denominator. Both the intensity ratio and the relative sensitivity Si/SIS bring in an analyte term and an IS term. For the IS element itself the ratios are identically 1 (fully correlated), so those terms cancel → its measurement σ is 0. The IS mass fraction CIS is excluded here (handled by the separate scaling layer).

6b. Total oxides (closure)

Ci=T·pi/D

pn=maxIn,0/Sn

D=Σnpn·cn

cn= oxide conversion factor ( cn=0 for non-oxide analytes)

T=total oxide wt·104

λ=T/D

σC,i2=λ2·Σnδin-Ci·cn/T2·σp,n2

σp,n2=σI,n/Sn2+pn·σS,n/Sn2

Explanation: there is no known internal standard; the absolute level is fixed by requiring all major-element oxides to sum to the specified total. Every analyte's uncertainty is funnelled through the closure sum D.

Reasoning: because the normalization λ=T/D depends on all majors, perturbing any element shifts every Ci. The Jacobian termδin-Ci·cn/T splits into the element's own response (δin) and the pull-back from the closure constraint -Ci·cn/T — this captures the constant-sum anti-correlation. Setting cn=0 for non-oxide analytes makes it one unified formula, and the result is invariant to which element is used internally as the reference.

6c. From-CAL internal standard

same 4-term quadrature as 6a, plus one inherited term:

σC,IS/CIS2=σsource/Csource2

Explanation: when the IS concentration is taken "from" another calibrator's result, that value is itself a measured quantity.

Reasoning: unlike a typed/known IS (whose uncertainty is external), an inherited IS carries a real, propagated measurement uncertainty from the source quantification. It is independent of this spot's own intensities and sensitivities (different acquisition), so it enters the quadrature as a fifth term. If the source is itself total-oxide or from-CAL, it resolves iteratively down the chain.

Final result: σC,i from Step 6 is the 1σ measurement uncertainty on the host concentration Ci, combining unknown-side counting (Steps 1–2) and the sensitivity's counting + RM (Steps 3–5), through the appropriate quantification model (Step 6).