# Uncertainty Propagation

ICP-Base · Technical Reference

How the 1σ measurement uncertainty of every mass fraction is calculated, step by step: first for matrix (host) spots, then for inclusions. Each step states what is computed, why it is done that way, and where it lives in the code.

# Overview and notation

### What is reported

For each analyte in each unknown spot, the software reports a mass fraction (in µg/g) and a 1σ measurement uncertainty. This uncertainty contains three kinds of sources: counting statistics of the signals, the calibration (sensitivities and the certified mass fractions of the reference materials), and the uncertainties that you enter yourself for the internal standard. It does not describe how well a model fits reality, for example a host that is not homogeneous.

### Three rules used throughout

- Products and quotients. When a result is a product or quotient of independent quantities, the relative variances add. This is the case for the direct internal standard calculation.

- Sums and differences. For a sum or difference of independent quantities, the absolute variances add. This is the case for background subtraction and for unmixing.

- General case. When quantities are reused or combined in a longer formula, the uncertainty of the result is obtained from the first-order rule below: the sum, over all inputs, of (partial derivative × input uncertainty) squared. Rules 1 and 2 are special cases of it.

General first-order rule for a result f that depends on independent inputs v

σf2=v∂f∂v2σv2

Errors from different measurements are independent and add in quadrature, and they shrink when averaged. Errors that come from one shared number, such as the certified mass fraction of a reference material, are correlated: they are the same in every spot that uses that material, and averaging does not reduce them. Keeping these two kinds apart is the main design principle of the calibration steps (A3 to A5).

### Data flow

The calculation runs in four stages. Steps A1 and A2 give the signal SEM of each spot. Steps A3 to A5 give the sensitivity and its uncertainty along the drift. Steps A6 to A8 give the matrix mass fractions. Steps B1 to B7 give the inclusion mass fractions. The signal SEMs and the calibration (A1 to A5) are computed in the drift module, the mass fractions and their uncertainties (A6 onward) in the quantification module. The inclusion calculation needs the host to be quantified first, because it uses the host mass fractions and their uncertainties.

### Switch and old sessions

The propagation in the quantification module is controlled by the menu Propagate uncertainties. If it is off, none of the uncertainty functions runs, and the results are unchanged. Sessions that were drift-corrected before this feature existed lack the stored SEM and sensitivity-uncertainty tables. The propagation then skips silently, and re-running the drift module creates the tables.

### Notation

| Symbol | Meaning |
| --- | --- |
| I | Background-corrected intensity in counts per second (CPS) |
| dt, N | Dwell time per sweep in seconds (per analyte), and number of sweeps in a window |
| σ | 1σ standard uncertainty. σX is the uncertainty of the quantity X |
| SEM | Standard error of a mean |
| C | Mass fraction, in µg/g (the certified mass fraction for a reference material, the result for an unknown) |
| S | Sensitivity, S=I/C (CPS per µg/g) |
| i, IS | The analyte of interest, and the internal standard (IS) |
| counting, RM | The two components of a sensitivity uncertainty: random counting error, and the (shared) uncertainty of the certified mass fraction |
| w | Position of an unknown between two calibration anchors, 0 at the left anchor and 1 at the right |
| p, c, D, T, λ | Quantities of the Total Oxides closure (A7) |
| mix, host, incl | Inclusion mode: the mixed signal (inclusion plus host), the pure host, and the unmixed inclusion |
| r, x | Inclusion mode: the host share of the mix (tracer ratio), and the inclusion share of the mix (mixing ratio) |
| m, H, hi | Inclusion mode: mass fraction of a standard in the inclusion, mass fraction of a standard in the host, and of the analyte in the host |

# Part A · Matrix

## A1 · Counting SEM of a window

calculate_background_corrected_matrix_intensities

### Explanation

For every spot, two time windows are used: the background window and the matrix window (up to three matrix intervals, which are pooled into one stack of sweeps). For each analyte and each window, the mean CPS is taken, and this step computes the standard error of that mean. It is done for all spots, calibrators and unknowns alike.

### Variables

| Symbol | Meaning |
| --- | --- |
| I | Mean CPS of the window (raw, not yet background-subtracted) |
| dt | Dwell time of the analyte, in seconds |
| N | Number of sweeps in the window |
| σPoisson | Counting-statistics SEM |
| σobserved | SEM from the actual sweep-to-sweep scatter |

### Formulas

Poisson SEM (physical lower bound)

σPoisson=Idt·N

Observed SEM (what the sweeps actually show)

σobserved=std(sweeps)N

SEM used (background and matrix window separately)

SEM=maxσPoisson, σobserved

### Reasoning

Poisson term. The instrument reports counts per second, so the number of counts registered in one sweep is I·dt. For a counting process the variance of the count equals the count, so the variance of the CPS of one sweep is I/dt. The mean of N independent sweeps has variance I/(dt·N). The dwell time is essential: leaving it out would underestimate the counting error by a factor of dt. This term is a hard physical floor, because no measurement can be more precise than counting statistics allow.

Observed term. The sweeps can scatter more than Poisson statistics predict, for example because the ablation is unsteady, the sample is heterogeneous, or the plasma fluctuates. The standard deviation of the sweeps divided by N captures this. It is the SEM and not the standard deviation, because every mass fraction uses the mean over the N sweeps.

Maximum, not sum. Both terms estimate the same thing, the uncertainty of the window mean. Adding them in quadrature would count the same scatter twice. Taking the larger value uses the better-informed estimate: the observed scatter when real extra noise is present, and the Poisson floor when the window is short and the observed value is itself noisy.

Edge cases. Spots without a background window or without any matrix window are skipped and keep SEM = 0. The pooled matrix window uses all sweeps of all its intervals as one stack.

## A2 · Background correction

calculate_background_corrected_matrix_intensities

### Explanation

The intensity used in all calculations is the matrix mean minus the background mean. This step gives the SEM of that net intensity. It is stored together with the net intensity for every spot.

### Formula

SEM of the net intensity

σnet=σwindow2+σbg2

### Reasoning

The net intensity is a difference. For a difference of independent quantities the absolute variances add. The two windows are measured at different times, so their counting errors are independent. Note that the background mean is the same number subtracted from every sweep of the matrix window, but it enters the result only once, as one subtraction of two means, which is exactly what is propagated here.

## A3 · Sensitivity of a CAL spot

calculate_sensitivities_of_CALs · group_CALs_and_bracket_ids

### Explanation

A calibrator (CAL) is a reference material with known certified mass fractions. For each CAL spot and analyte, the sensitivity is the net intensity divided by the certified mass fraction. Its uncertainty is split into two components that are carried separately from here on.

### Variables

| Symbol | Meaning |
| --- | --- |
| I, σI | Net intensity of the CAL spot and its SEM from A2 |
| C, σC | Certified mass fraction of the analyte in the reference material, and its 1σ uncertainty from the reference material file (0 if none is given) |

### Formulas

Sensitivity

S=IC

The two uncertainty components

σS,counting=S·σII

σS,RM=S·σCC

### Reasoning

The sensitivity is a quotient, so its relative uncertainty combines the relative uncertainties of the intensity and of the certified mass fraction. They are kept as two components because they behave differently in the next steps:

- Counting. Random and independent for each CAL spot. It shrinks when several CAL spots are averaged.

- RM. Systematic. Every CAL spot of the same material uses the same certified value, so this error is identical in all of them (fully correlated). It does not shrink with averaging, and it is the same at both anchors of a drift segment.

If the two were merged here, the averaging and the interpolation below would treat the systematic part as random and wrongly shrink it. The per-spot combined value (the quadrature sum of the two) is only used for the error bars of the single CAL points in the drift plot.

## A4 · Averaging a group of CAL spots

group_CALs_and_bracket_ids

### Explanation

Consecutive CAL spots of the same reference material, with no other spot type in between, form a group. The group is averaged into one sensitivity, which serves as a calibration anchor at the middle of the group. This step gives the uncertainty of that anchor.

### Variables

| Symbol | Meaning |
| --- | --- |
| Sj | Sensitivity of the j-th CAL spot of the group |
| σcps,j | SEM of the net intensity of the j-th spot (A2) |
| N | Number of CAL spots in the group (often 2 or 3) |
| S | Mean sensitivity of the group |

### Formulas

Group mean

S=1NjSj

Counting part: propagated floor, observed scatter between the CAL spots, and the larger of the two

σfloor=jσcps,j2N·C

σbetween=stdS1, S2, …, SNN

σcounting=maxσbetween, σfloor

RM part (correlated, does not shrink with N)

σRM=S·σCC

Combined sensitivity uncertainty of the anchor

σS=σcounting2+σRM2

### Reasoning

Floor. The mean of N independent measurements has the uncertainty σ2/N in CPS, which is divided by the certified mass fraction to give the sensitivity. This is what counting statistics alone predict, and it shrinks as 1/N.

Observed scatter. If the CAL spots disagree more than counting statistics predict (inhomogeneity of the standard, fast drift inside the group), the standard error of the sensitivities captures that disagreement directly.

Maximum, not sum. Both estimate the same uncertainty of the group mean, and when the standard is homogeneous the scatter between spots is itself caused by counting noise. Adding them in quadrature would count it twice. With the maximum, agreeing CAL spots give the counting floor, and disagreeing spots give the observed scatter. For a single CAL spot the standard deviation is 0, so the floor is used automatically.

Why not a χ² test (MSWD). With only 2 or 3 spots per group the test has 1 or 2 degrees of freedom, which is far too little to estimate an inflation factor reliably. The maximum is simpler, transparent, and robust at small N.

RM part. The certified mass fraction is one number shared by all spots of the group, so its uncertainty is fully correlated and enters once, as a relative uncertainty on the mean sensitivity. It does not shrink with N.

## A5 · Drift interpolation to the unknowns

linear_regression · unknown_fitting

### Explanation

The group anchors are discrete points in time. An unknown spot between two anchors A and B gets its sensitivity from the straight line through the two anchors, evaluated at the spot's time. The uncertainty is interpolated too, and the two components are treated differently. The sensitivity of the CAL spots themselves is not needed for unknowns, so their rows are set to zero.

### Formulas

Position and interpolated sensitivity (t = time of the unknown)

w=t−tAtB−tA

S(t)=(1−w) SA+w SB

Counting: independent anchors, quadrature (the band narrows between the anchors)

σcounting(w)=1−w2σc,A2+w2 σc,B2

RM: correlated anchors, linear (carried through in full)

σRM(w)=(1−w) σRM,A+w σRM,B

Combined

σS=σcounting2+σRM2

### Reasoning

A point on the line is a weighted sum of its two anchors. If the anchors carry independent errors, the weighted sum combines in quadrature, and the uncertainty is smallest in the middle between the anchors. If both anchors carry the same error, as the shared certified mass fraction does, the weights simply add up to one and the error is carried through in full, so it does not shrink between the anchors. Treating the RM part with the quadrature rule would wrongly make it smaller in the middle. This is the same independent-versus-correlated distinction as in A4, now along the time axis.

Both the combined value and the RM part are stored for every unknown, so that later steps can use the combined value and the RM part can be separated again if needed.

## A6 · Mass fraction with a direct internal standard

calculate_matrix_uncertainties_internal_standard

### Explanation

When the internal standard (IS) mass fraction is entered by you, the mass fraction of analyte i is the ratio of the analyte and IS intensities, multiplied by the IS mass fraction and corrected by the sensitivity of the analyte relative to the IS. This step gives its uncertainty.

### Variables

| Symbol | Meaning |
| --- | --- |
| Ii, IIS | Net intensities of the analyte and the IS in the unknown spot, with SEM σI from A2 |
| Si, SIS | Interpolated sensitivities of the analyte and of the IS element (A5), with the combined uncertainty σS |
| CIS, σC,IS | IS mass fraction that you entered, and the ±1σ that you entered (0 if left empty) |

### Formulas

Mass fraction

Ci=IiIIS·CISSi/SIS

The five relative variance terms

V1=σI,iIi2  analyte counting

V2=σS,iSi2  analyte sensitivity

V3=σI,ISIIS2  IS counting

V4=σS,ISSIS2  IS sensitivity

V5=σC,ISCIS2  your ± on the IS mass fraction

Uncertainty of the mass fraction

σi=CiV1+V2+V3+V4+V5

### Reasoning

The mass fraction is a pure product and quotient of independent quantities, so the relative variances add (rule 1). Each factor contributes whether it sits in the numerator or the denominator. Both the intensity ratio and the relative sensitivity contain an analyte term and an IS term, which gives the four measurement terms. The fifth term is only added if you entered an uncertainty on the IS mass fraction.

The IS column. For the internal standard element itself, the intensity ratio and the sensitivity ratio are both exactly 1: the same signal appears in numerator and denominator. Terms V1 to V4 are therefore set to 0 for that column. Only V5 remains, so its uncertainty equals the ± you entered (0 if none). For every other analyte, V3 and V4 stay in, because the IS intensity is the denominator of every ratio.

Sensitivity uncertainty. V2 and V4 use the combined interpolated sensitivity uncertainty (counting and RM together). Quantification treats it as one number per element.

Not covered here. Spots whose IS mass fraction comes from another calibrator are handled in A8, and Total Oxides spots in A7. Spots without a valid IS or with all-zero intensities keep σ = 0.

## A7 · Mass fraction with Total Oxides

calculate_matrix_uncertainties_total_oxides

### Explanation

With Total Oxides there is no IS mass fraction. The absolute level is fixed by requiring the oxides of all major elements to add up to a total that you specify (for example 100 wt%). Every mass fraction is therefore scaled by one common factor that depends on all majors. A change in any one element's signal shifts every mass fraction, and this closure creates anti-correlations that must be propagated with the full partial derivatives (rule 3).

### Variables

| Symbol | Meaning |
| --- | --- |
| pn | Preliminary mass fraction of element n: max(In,0)/Sn |
| σp,n | Uncertainty of pn, from the SEM of In and the uncertainty of Sn |
| cn | Effective oxide factor of element n (mass of oxide per mass of element). 0 for analytes that do not form an oxide in the sum. Iron is split into FeO and Fe₂O₃ using the Fe ratio you entered |
| D | Closure sum, D=npncn over the oxide-forming majors |
| ox, σox | Total oxide sum that you entered (wt%) and your ±1σ on it |
| T | The same total in µg/g: T=104·ox |
| λ | Normalisation factor, λ=T/D |
| δin | Kronecker delta: 1 if n is analyte i, otherwise 0 |

### Formulas

Preliminary mass fractions, closure sum and result

pn=max(In,0)Sn

D=npn cn

λ=TD

Ci=λ pi

Effective oxide factor of iron (a = FeO share from the Fe ratio f; k = 0.699480 / 0.777314)

a=11+k 1−ff

cFe=a cFeO+(1−a)·1.4297

Derivative of the result with respect to each preliminary mass fraction

∂Ci∂pn=T δinD−T pi cnD2=λδin−Ci cnT

Uncertainty of each preliminary mass fraction

σp,n2=σI,nSn2+pn σS,nSn2

Closure uncertainty (sum over the oxide-forming majors)

σclosure,i2=λ2nδin−Ci cnT2σp,n2

An analyte that is not an oxide-forming major adds its own term once

σclosure,i2→σclosure,i2+λ2 σp,i2

Your uncertainty on the oxide sum (relative term, added last)

σi=σclosure,i2+Ci σoxox2

### Reasoning

Closure. With Ci=T pi/D, the result depends on pi directly and on every pn through D. Differentiating gives two parts: the element's own response (the δ term) and the pull-back from the constant sum (the Cicn/T term). Together they give the anti-correlation that makes the uncertainty of a major element smaller than its raw counting error would suggest.

Non-oxide analytes. An analyte that is not in the sum (c=0, not a major) is scaled by λ but does not enter D. Its response to the majors is the pull-back term only, and its own σp is added once. This single formula holds for majors and for all other elements, with no separate treatment.

Clamping. p uses max(I,0) because a negative signal cannot represent a negative oxide in the sum. In σp, the counting part (σI/S) is kept even when I≤0, so the uncertainty does not collapse to zero for a signal below background.

Your oxide ±. The uncertainty of the total oxide sum scales every mass fraction by the same relative amount. It is independent of the measurement, so it is added in quadrature at the end.

## A8 · Internal standard taken from another calibrator

calculate_matrix_uncertainties_from_cal_passes

### Explanation

The IS mass fraction of a spot can be set to "from <another calibrator>". The spot is then quantified against calibrator A, but the IS mass fraction is the result for the same spot quantified against calibrator B. That value is itself a measured quantity with its own uncertainty, which has to enter this calculation.

Same four measurement terms as A6, plus the inherited term

σi=CiV1+V2+V3+V4+σsourceCsource2

### Reasoning

Unlike an IS mass fraction that you type in (whose uncertainty comes from your ±), an inherited one carries a real propagated measurement uncertainty. It comes from the source quantification, so it is independent of this spot's own intensities and sensitivities (a different calibration), and it enters as a fifth independent term. Your ± entry for such a cell is discarded in the table, because the inherited uncertainty takes its place.

Order and iteration. The source may itself be a Total Oxides result or another "from" spot. For that reason this function runs last, and it repeats until no value changes, with at most as many passes as there are calibrators. Chains of "from" spots then settle one link per pass.

## A9 · Inputs and conventions (matrix)

### Your inputs

With the menu Enable internal standard uncertainty checked, the matrix table gets two ± columns. Both are absolute ±1σ in the unit of the column.

| Column | Used for |
| --- | --- |
| ± 1σ [µg/g] | The IS mass fraction (term V5 in A6). Ignored for Total Oxides spots and for values taken from another calibrator |
| ± 1σ [%] | The total oxide sum (last step of A7) |

When the IS is chosen for a QCRM spot, the IS mass fraction and its ± are filled in from the certified values of that quality-control material (if it has an uncertainty). Changing the IS analyte clears the ±.

### Conventions

- Empty ±. Means "this input is exact" and adds nothing. Everything measured (counting, sensitivities and the certified mass fractions of the calibrators) is still propagated. The result then covers the measurement uncertainty only.

- Negative values. Nothing is clipped. A negative mass fraction is below the detection limit and not meaningful. Its uncertainty is calculated by the same formula (σ = mass fraction × relative uncertainty), so it carries the sign of the mass fraction. Such values should be disregarded.

- Independence. Analytes are treated as independent. In reality they share the IS signal and the calibration, which makes some errors correlated. The IS terms enter every analyte fully, so shared parts partly cancel in reality; ignoring this errs on the safe side.

- Not yet included. The uncertainty of the interference correction and of the laser-blank correction. The SEMs contain counting statistics only.

# Part B · Inclusion

## B0 · What changes for inclusions

### Explanation

An inclusion is measured together with the host around it. Its signal window therefore contains a mixture of inclusion and host. Everything for the host stays as in Part A, including the calibration (A1 to A5): inclusions use the same sensitivities and sensitivity uncertainties. What is new is how the mixed signal is turned into an inclusion mass fraction. There are three ways, set per spot:

| Mode | How the host share is found | Steps |
| --- | --- | --- |
| Matrix-only tracer (MOT) | An element that exists only in the host. Its signal in the mix, relative to a pure host window, gives the host share. The host signal is subtracted, and the inclusion is quantified from the remaining CPS (direct standard or Total Oxides). | B1–B4 |
| Mixing ratio | You enter the inclusion share x of the mix. The inclusion mass fraction is derived from the mixed mass fraction minus the host. | B5 |
| Two internal standards | Two standards with known mass fractions in the inclusion and in the host determine x from the mix. | B6 |

### Reasoning

The three modes have different mathematical forms. MOT keeps the product-and-quotient form of Part A after an unmixing subtraction, so it can be propagated with relative variances. The other two contain sums, differences and a shared mixing ratio, so they need explicit partial derivatives (rule 3). Each spot uses exactly one mode, and the modes do not feed into each other.

The sensitivity uncertainty, the host mass fraction and its uncertainty, and your ± entries are inputs. Because the inclusion calculation uses the host mass fraction and its σ, the host spot has to be quantified with the uncertainty propagation switched on.

## B1 · SEM of the mixed signal

calculate_background_corrected_mix_intensities

### Explanation

The inclusion window gets its own SEM, exactly as in A1 and A2: the larger of Poisson and observed SEM on the raw inclusion window, combined in quadrature with the SEM of the spot's own background window.

SEM of the mixed signal of analyte i

σmix,i=SEMincl. window,i2+SEMbg,i2

### Reasoning

The mixed signal is a background-corrected mean, like any other signal, so it is handled the same way. The background SEM is computed here from the inclusion spot's own background window. It is not taken from the matrix SEM table, which is skipped for spots without matrix intervals and would then silently be zero. The result is stored for all non-calibrator spots that have an inclusion interval.

## B2 · Matrix-only tracer unmixing

matrix_only_tracer_unmixing (drift module)

### Explanation

The tracer is an element that exists only in the host. Its signal in the mix, compared with its signal in a pure host window, tells how much host is in the mix: r=Mmix/Mhost. The host signal, scaled by r, is subtracted from every analyte. What is left is the inclusion CPS. The host window comes from the host spot, which can be the spot itself or another one. It is calculated after any signal corrections, in the drift module, and the version in the interval module is only a preview.

### Variables

| Symbol | Meaning |
| --- | --- |
| mixi, σmix,i | Net CPS of analyte i in the inclusion window and its SEM (B1) |
| hi, σh,i | Net CPS of analyte i in the host spot (all its matrix windows pooled) and its SEM (A2) |
| Mmix, Mhost | The tracer signal in the mix and in the host, with SEMs σM,mix and σM,host |
| r, Vr | Host share of the mix, and its relative variance |
| incli | Inclusion CPS of analyte i after unmixing |

### Formulas

Host share and unmixing

r=MmixMhost

incli=mixi−hi r

Relative variance of the host share (a quotient of two independent signals)

Vr=σM,mixMmix2+σM,hostMhost2

Uncertainty of the inclusion CPS

σ2(incli)=σmix,i2+r σh,i2+hi r2Vr

- Mix counting. σmix,i2: counting error of the inclusion window, including its background.

- Host counting. (r σh,i)2: counting error of the host signal, scaled down by the host share.

- Tracer ratio. (hi r)2Vr: the error of the host share acts on the whole host contribution. The larger the host share, the more it matters.

### Reasoning

Where the formulas come from. The mix is inclusion plus host. A fraction r of the pure host window is present in the mix, so the host contributes hi r to analyte i. The tracer exists only in the host, so its whole signal in the mix is host signal, Mmix=r Mhost, and this fixes r. The inclusion CPS then depends on three inputs (mixi, hi and r), and r in turn on the two tracer signals. First-order propagation (rule 3) needs these derivatives:

| Input | Derivative |
| --- | --- |
| mixi | ∂ incli∂ mixi=1 |
| hi | ∂ incli∂ hi=−r |
| r | ∂ incli∂ r=−hi |
| Mmix (through r) | ∂r∂Mmix=1Mhost=rMmix |
| Mhost (through r) | ∂r∂Mhost=−MmixMhost2=−rMhost |

Putting it together

σr2=r2 Vr=r2σM,mixMmix2+σM,hostMhost2

σ2(incli)=σmix,i2+r2σh,i2+hi2 σr2=σmix,i2+r σh,i2+hi r2Vr

In short, the inclusion CPS is a difference between the mixed signal and a scaled host signal. For a difference the variances add (rule 2). The scaled host term is a product of the host signal and r, so the relative variances of the host signal and of r add; written out for the product this gives the host-counting and tracer-ratio terms. The three terms are independent, because the mix, the host and the tracer comparison come from different time windows.

The tracer itself. For the tracer, h·r equals its mix signal exactly, so its inclusion CPS is 0 and its uncertainty is 0. The general formula would count its error twice, so the tracer is set explicitly: incl = 0 and σ = 0. This does not remove the need for a positive tracer signal in the mix and in the host (see below).

Approximation. The shared r links the errors of all analytes, but the analytes are treated as independent. In a mass fraction, which is a ratio of two inclusion intensities, the shared part would partly cancel, so ignoring it errs on the safe side. If the host is the same spot as the inclusion, the two also share one background, which is ignored in the same way.

Tracer without a usable signal. r only makes sense if the tracer signal is positive in both the mix and the host. If either one is zero or negative after background, interference and blank correction, the drift module shows the alert "Unmixing Failed" for that spot, but the calculation carries on. The inclusion CPS and their uncertainties of that spot are then meaningless and should be discarded.

Negative values. For analytes that live mostly in the host, the subtraction removes most of the signal, and counting noise decides the sign. Negative inclusion CPS are therefore normal there. They are kept as they are, and the uncertainty is still a positive number.

## B3 · MOT, direct internal standard

calculate_inclusion_uncertainties_MOT

### Explanation

After unmixing, the inclusion CPS are quantified in the same way as the matrix (A6), with the inclusion CPS and their SEM from B2 in place of the matrix CPS and SEM. The sensitivities and their uncertainty are the same ones.

Mass fraction

Ci=incliinclIS·CISSi/SIS

Uncertainty: the four measurement terms plus your ± on the IS mass fraction (V2, V4 and V5 as in A6)

σi=CiV1+V2+V3+V4+V5

V1=σincl,iincli2

V3=σincl,ISinclIS2

### Reasoning

The structure is a product and quotient, so relative variances add (rule 1), as in A6. The only difference is that the counting term of each intensity is now the inclusion SEM from B2, which already contains the mix, host and tracer-ratio errors.

- IS column. The measurement terms cancel, and the uncertainty equals your ± on the IS mass fraction (0 if empty).

- Tracer column. Mass fraction 0 and uncertainty 0 by construction (B2).

- From-CAL standard. Handled in B7. Such spots are skipped here.

## B4 · MOT, Total Oxides

calculate_inclusion_uncertainties_MOT_total_oxides

### Explanation

For MOT spots whose internal standard is Total Oxides, the closure propagation of A7 is used unchanged, with inclusion inputs. The preliminary mass fraction of each major element is p=max(incl,0)/S, with the inclusion CPS from B2, the sensitivities of Part A, and the target total from the inclusion table.

Uncertainty of a preliminary mass fraction (same form as A7, with the inclusion SEM)

σp,n2=σincl,nSn2+pn σS,nSn2

Closure and your ± on the oxide sum (as in A7)

σclosure,i2=λ2nδin−Ci cnT2σp,n2

σi=σclosure,i2+Ci σoxox2

### Reasoning

The closure mathematics does not depend on where the preliminary mass fractions come from, so the derivation of A7 applies directly. The only change is that the counting error of each major is the inclusion SEM, which carries the mix, host and tracer-ratio parts. The tracer, if it is one of the majors, has p=0 and σp=0 and therefore contributes nothing. The closure inputs (the list of majors, the oxide factors, D and T) are stored while the mass fractions are calculated and then reused here.

### Limitation: majors that the host shares with the inclusion

The closure needs the inclusion value of every oxide-forming major, including the elements that the host and the inclusion have in common. With MOT these are the weakest numbers of the whole calculation: the inclusion signal of such an element is a small difference of two large ones. For an olivine-hosted inclusion, Mg, Si and Fe can come out near zero or negative, with an SEM several times larger than the value. Their uncertainty then enters the closure sum D, and through λ=T/D it enters every analyte.

Relative variance of an analyte that is not an oxide-forming major

σclosure,iCi2=ncn σp,nD2+σp,ipi2

The sum over the majors is the same for every analyte, so one badly determined major raises the relative uncertainty of all analytes by the same amount. The second term is the analyte's own counting error and is what a single internal standard would show.

Observed example. Olivine-hosted inclusions with Mn and Co as tracers. The terms c σp/D of Mg, Si and Fe were 5.7, 4.2 and 1.0 in one spot (and 6.2, 4.7, 1.1 in another), in the ratio of the oxides in olivine (MgO : SiO₂ = 1.33, forsterite 1.34). The error is therefore proportional to the amount of host. The inclusion values were 1.1·10⁶ ± 8.3·10⁶ cps for Mg, −1.4·10⁵ ± 8.2·10⁵ cps for Si and −1.2·10³ ± 5.2·10⁵ cps for Fe. The reported relative uncertainty was about 700% for all analytes, while the same spots with a single internal standard gave 30 to 50%. The code was checked for this case: the matrix and inclusion functions are identical in form, and the tracer contributes 0.

How to read the number. When the closure term is well above 1, the first-order propagation no longer describes the real scatter, because the clamp max(I,0) makes the distribution one-sided. Read such a value as "the normalisation is not determined", not as a precise percentage. The point values are affected too: Si and Fe are clamped to zero, D is then far too small, and λ scales all other elements up.

When to use it. Total Oxides with MOT is only meaningful when the host does not share its major elements with the inclusion. For hosts such as olivine, use an internal standard, the mixing ratio mode or the two standards mode, where the host majors do not enter. The software does not warn about this case at present.

## B5 · Mixing ratio mode

calculate_inclusion_uncertainties_mixing_ratio

### Explanation

Here you enter x, the share of the inclusion in the mix (0<x≤1). The mixed signal is first converted to a mixed mass fraction with the IS, and the host share (1−x) is then subtracted to give the inclusion.

### Variables

| Symbol | Meaning |
| --- | --- |
| Ii, IIS | Net CPS of the analyte and of the IS in the inclusion window (mix), with SEM from B1 |
| Si, SIS | Interpolated sensitivities (A5) with their combined uncertainty |
| m, σm | Mass fraction of the IS in the inclusion (entered) and your ±1σ on it |
| H, σH | Mass fraction of the IS element in the host and its uncertainty (matrix result, Part A) |
| hi, σh,i | Mass fraction of the analyte in the host and its uncertainty (matrix result) |
| x, σx | Mixing ratio that you entered, and your ±1σ on it |

### Formulas

Mixed mass fraction of the standard, and the analyte factor

M=(1−x) H+x m

Ki=IiIIS·SISSi

Inclusion mass fraction

Cincl,i=Ki Mx−1−xx hi

First-order propagation, all inputs independent

σC2=v∂Cincl,i∂v2σv2

### Derivation of the inclusion mass fraction

1 · Mass balance. The mix consists of a share x of inclusion and (1−x) of host, so every element j obeys

Cmix,j=x Cincl,j+(1−x) Chost,j

2 · Quantify the mix. The mix is quantified like an unknown against the standard (as in A6). The mass fraction of the standard in the mix is not known directly, but it follows from the mass balance of step 1 applied to the standard itself:

Cmix,i=Ki M

M=Cmix,IS=(1−x) H+x m

3 · Solve for the inclusion. Step 1 solved for the inclusion, with step 2 inserted, gives the formula above:

Cincl,i=Cmix,i−(1−x) hix=Ki Mx−1−xx hi=1x−1Ki H−hi+Ki m

Assumptions. The sensitivities of the calibrator also hold for the mix (no extra matrix effect between host and inclusion). The mixing is by mass. The host in the mix has the composition of the quantified host spot, which is why the host mass fractions and their uncertainties are inputs of the propagation.

Where the derivatives come from. The four intensity and sensitivity inputs act only through Ki. Since ∂C/∂K=M/x and K is a product and quotient, ∂K/∂Ii=K/Ii, ∂K/∂Si=−K/Si, ∂K/∂IIS=−K/IIS and ∂K/∂SIS=K/SIS, the chain rule gives the first four rows of the table. H, hi, m and x act directly, and the last form of the formula above shows their derivatives at a glance: ∂C/∂H=K (1−x)/x, ∂C/∂hi=−(1−x)/x, ∂C/∂m=K, and ∂C/∂x=−(K H−hi)/x2.

### Partial derivatives for each input

| Input v | ∂Cincl,i/∂v |
| --- | --- |
| Ii (mix CPS of the analyte) | Mx·SISIIS Si |
| Si (sensitivity of the analyte) | −Mx·KiSi |
| IIS (mix CPS of the IS) | −Mx·KiIIS |
| SIS (sensitivity of the IS) | Mx·KiSIS |
| H (host, IS element) | Ki 1−xx |
| hi (host, analyte) | −1−xx |
| m (IS mass fraction) | Ki |
| x (mixing ratio) | hi−Ki Hx2 |

### Reasoning

Why partial derivatives. The formula contains a difference, and several inputs (x, H, the CPS and sensitivities through Ki) appear in more than one place. Relative variances do not simply add in such an expression, but the first-order rule always applies: each input contributes (derivative × uncertainty) squared, and independent inputs add. All derivatives are written out explicitly in the code.

Derivative with respect to x. Since C=(K H−h)(1/x−1)+K m, the derivative with respect to x is −(K H−h)/x2=(h−K H)/x2. This is why an uncertainty in x grows like 1/x2.

The standard itself. By construction, the inclusion value of the standard equals its own mass fraction m. Its uncertainty is therefore just your ± on m (0 if empty), and the formula is not used for that column.

Small x amplifies errors. The inclusion is obtained by removing a large host contribution, so errors of the mixed signal and of the host mass fraction are magnified by about 1/x: roughly 2 at x=0.5, 10 at x=0.1 and 100 at x=0.01. This is physical, and it means that very small inclusions are poorly constrained.

Dependencies. The host spot must be quantified with its uncertainty. If you leave the ± on x empty, x is treated as exact. That is the least complete case, because x enters through 1/x.

Independence. The sum over inputs assumes that all inputs are independent. This is not fully true. The host mass fractions H and hi were obtained with the same calibration as the sensitivities Si and SIS used here (the same reference materials, drift fit and, for the host, possibly the same IS), so part of their errors is shared. The covariance between them is not modelled. Depending on the signs it can make the true uncertainty smaller or larger than the one reported, and the direction has not been determined. The same assumption applies to the matrix calculation (A9) and to the other inclusion modes.

## B6 · Two internal standards mode

calculate_inclusion_uncertainties_two_IS

### Explanation

Two standards (IS1 and IS2) have known mass fractions m1 and m2 in the inclusion and are measured in the host (H1, H2). The ratio of their signals in the mix determines the mixing ratio x. After that the calculation continues as in B5, using the second standard as the reference.

### Formulas

Ratio of the two standards in the mix, and the mixing ratio

a=I1I2·S2S1

x=H1−a H2H1−m1−a H2+a m2

Inclusion mass fraction (reference: the second standard)

M2=(1−x) H2+x m2

Ki=IiI2·S2Si

Cincl,i=Ki M2x−1−xx hi

Partial derivatives of x (N = H₁ − a H₂, D = H₁ − m₁ − a H₂ + a m₂)

∂x∂H1=a m2−m1D2

∂x∂H2=a (m1−a m2)D2

∂x∂m1=ND2

∂x∂m2=−a ND2

∂x∂a=H2 (m1−a m2)−N m2D2

Partial derivatives of a

∂a∂I1=aI1

∂a∂I2=−aI2

∂a∂S1=−aS1

∂a∂S2=aS2

Total derivative with respect to any input v: direct part plus the part through x

dCincl,idv=∂Cincl,i∂vx+∂Cincl,i∂x·dxdv

σC2=vdCincl,idv2σv2

### Reasoning

Derivation of x. Both standards are measured in the mix. Quantified like unknowns (C=I/S), the ratio of their mass fractions in the mix is the ratio of their sensitivity-corrected intensities, and each of them obeys the mass balance of B5:

Ratio of the standards in the mix

a=Cmix,1Cmix,2=I1/S1I2/S2=I1I2·S2S1

Mass balance of each standard

Cmix,1=(1−x) H1+x m1

Cmix,2=(1−x) H2+x m2

Insert and solve for x

(1−x) H1+x m1=a(1−x) H2+x m2 ⟹ H1−a H2=xH1−m1−a H2+a m2

The last line is the expression for x above, with N the left side and D the bracket. Its derivatives follow from the quotient rule, ∂x/∂v=(∂N/∂v·D−N·∂D/∂v)/D2, and from D−N=a m2−m1. For example, N and D both rise by 1 with H1, so ∂x/∂H1=(D−N)/D2=(a m2−m1)/D2. With m1, N does not change and D falls by 1, so ∂x/∂m1=N/D2. With H2 both fall by a, so ∂x/∂H2=a (N−D)/D2=a (m1−a m2)/D2. With m2, D rises by a, so ∂x/∂m2=−a N/D2. With a, N changes by −H2 and D by m2−H2, so ∂x/∂a=(H2 (m1−a m2)−N m2)/D2. The derivatives of a with respect to the intensities and sensitivities are those of a product and quotient, as listed.

Once x is known, the analyte follows exactly as in B5, with the second standard as the reference (Cmix,2=M2). The first standard is only used to find x.

Consequence. By construction, x is the value for which the mixed amount of IS1, a M2, equals (1−x) H1+x m1. This guarantees that the inclusion value of each standard equals its own mass fraction (Cincl,1=m1 and Cincl,2=m2). Their uncertainty is therefore your ± on that mass fraction (0 if empty), and the derivative formulas are used for all other columns.

Where each input acts. Ii and Si enter only through Ki; hi enters directly; I1 and S1 enter only through a, and therefore only through x; I2 and S2 enter through Ki directly and through a; H1 and m1 enter only through x; H2 and m2 enter directly (through M2) and through x. The total derivative adds the direct part and the part through x for each of them.

Independence and amplification. The assumption that all inputs are independent (the host mass fractions and the inclusion share the same calibration) is explained in the Independence note of B5, and the amplification of errors at small x in the note "Small x amplifies errors" of B5. Both apply here too. In this mode there is one more shared quantity: I2 and S2 act on the result both directly and through x, which the total derivative accounts for, whereas the correlation between host and inclusion is ignored.

## B7 · Standards taken from another calibrator

calculate_inclusion_uncertainties_from_cal_passes

### Explanation

As in A8, the mass fraction of IS1 (and in two-IS mode also of IS2) can be set to "from <another calibrator>". That value is a result with its own uncertainty. The three mode functions (B3 or B4, B5, B6) are written for a direct value with a user-entered ±, so instead of copying their mathematics, the following cycle is used:

- Look up. Read the inherited mass fraction and its σ from the source calibrator's mass fraction and uncertainty tables (same spot, same standard).

- Substitute. Write both into the input fields in place of the "from" entry, as if you had entered a value with a ±.

- Recompute. Run the mode functions, then restore the original entries.

- Repeat. Repeat until nothing changes, so a source that is itself "from CAL" settles. The maximum is the number of calibrators.

### Reasoning

Mathematically this is the same as A8: the inherited σ enters as the uncertainty of the standard's mass fraction m. In the product form (B3) that is the fifth term (σm/m)2, and in the longer formulas (B5, B6) it is the input uncertainty σm in the partial-derivative sum. The inherited value is independent of this spot's own intensities, because it comes from a different calibration.

Skipped when not needed. If no spot takes a mass fraction from another calibrator, the function returns immediately, so nothing is recomputed. A spot whose source mass fraction is 0 is left untouched, as in the mass fraction calculation itself.

## B8 · Inputs and conventions (inclusion)

### Your inputs

With the menu Enable internal standard uncertainty checked, the inclusion table gets ± columns. All are absolute ±1σ in the unit of the column.

| Column | Used for |
| --- | --- |
| ± 1σ [µg/g] (I.S.1) | The first standard's mass fraction (B3, B5, B6). Discarded if the value comes from another calibrator or the standard is Total Oxides |
| ± 1σ [%] | The total oxide sum (B4) |
| ± 1σ [µg/g] (I.S.2) | The second standard, in two-IS mode only (B6) |
| ± 1σ (x) | The mixing ratio, in Mixing ratio mode only (B5) |

There is no automatic filling of ± values for inclusions. Changing the analyte of IS1 clears the ± that no longer belongs to it.

### Conventions

- Empty ±. Means "this input is exact". Everything measured (counting, sensitivities, host mass fractions with their uncertainty) is still propagated. The result then covers measurement uncertainty only, which is least complete in Mixing ratio mode, where x enters through 1/x.

- Negative values. Nothing is clipped. For analytes that live mostly in the host, the unmixed value can be negative. It is kept as it is, and the uncertainty of the inclusion is calculated with the formulas above (for MOT it keeps the sign of the mass fraction, as in the matrix).

- Independence. All inputs are treated as independent. Shared sensitivities between host and inclusion, and the shared host share r in MOT, are not modelled. For MOT this errs on the safe side.

- Dependency on the host. Mixing ratio and two-IS modes need the host spot to be quantified with uncertainty. If it is not, the affected spots keep σ = 0.

- Not yet included. The uncertainty of the interference correction and of the laser-blank correction. The SEMs contain counting statistics only. There is also no uncertainty on the mixing ratio or the mixed mass fraction as outputs.